## ACR metrics
| ACR Test | Subtest slice # | Subtest | IRSE Pulseq | IRSE Siemens
1 | Geometric accuracy | 1 | Top-to-bottom diameter | 190.4395 | 190.4395
|  | 1 | Left-to-right diameter | 189.6973 | 191.1621
|  | 5 | Top-to-bottom diameter | 191.4258 | 191.9043
|  | 5 | Left-to-right diameter | 189.707 | 191.1621
|  | 5 | NE-SW diameter | 190.4199 | 191.7969
|  | 5 | NW-SE diameter | 190.9473 | 190.9668
|  |  | Passed? (all within 2 mm of 190 mm) | YES | YES
2 | High-contrast spatial resolution | 1 | Best row resolution | 0.9 | 0.9
|  | 1 | Best column resolution | 0.9 | 0.9
|  |  | Passed? (resolution at least 1 mm) | YES | YES
3 | Slice thickness accuracy | 1 | Ramp length (top) | 41.0156 | 41.5039
|  | 1 | Ramp length (bottom) | 46.7578 | 50.75
|  | Derived | Slice thickness | 4.3699 | 4.5664
|  |  | Passed? Within 5 mm +- 1 mm | YES | YES
4 | Slice position accuracy | 1 | Bar length difference (mm) | -1.0352 | -3.3008
|  | 11 | Bar length difference (mm) | -3.5449 | 5.7422
|  |  | Passed? < 5 optimal, 5-7 ok | YES | YES
5 | Image intensity uniformity | 7 | Large ROI low-signal value | 0.419 | 0.3023
|  | 7 | Large ROI high-signal value | 0.6471 | 0.4956
|  | Derived | Percent integral uniformity (PIU) | 78.6043 | 75.7739
|  |  | Passed? > 82% optimal, <80% fail | NO | NO
6 | Percent-signal ghosting | 7 | Large ROI mean pixel intensity | 0.5023 | 0.3723
|  | 7 | Left ROI MPI | 0.0054 | 0.0121
|  | 7 | Right ROI MPI | 0.0071 | 0.0113
|  | 7 | Top ROI MPI | 0.0142 | 0.0071
|  | 7 | Bottom ROI MPI | 0.0148 | 0.0074
|  | Derived | Ghosting ratio | 0.016424 | 0.011953
|  |  | Passed? < 0.025; fail if > 0.030 | YES | YES
7 | Low-contrast object detectability | 8 | # Complete spokes | 8 | 3
|  | 9 | # Complete spokes | 8 | 10
|  | 10 | # Complete spokes | 8 | 10
|  | 11 | # Complete spokes | 10 | 10
|  | Derived | Total score | 34 | 33
|  |  | Passed? >37 | NO | NO

## Derived or display values
Name | IRSE Pulseq | IRSE Siemens | TSE Pulseq | TSE Siemens
Test 1 display center | 0.25 | 0.19705 | 0.23 | 0.23
Test 1 display width | 0.5 | 0.3941 | 0.46 | 0.46
Test 3 top mean | 0.0657 |  | 0.0653 | 0.04
Test 3 bottom mean | 0.0658 |  | 0.069 | 0.0418
Test 3 mean | 0.06575 | 0.0383 | 0.06715 | 0.0409
0.032875