## ACR metrics
| ACR Test | Subtest slice # | Subtest | TSE Pulseq | TSE Siemens
1 | Geometric accuracy | 1 | Top-to-bottom diameter | 189.9 | 188.7
|  | 1 | Left-to-right diameter | 190.7 | 191.9
|  | 5 | Top-to-bottom diameter | 191.4 | 189.98
|  | 5 | Left-to-right diameter | 190.4 | 191
|  | 5 | NE-SW diameter | 191.2 | 190.7
|  | 5 | NW-SE diameter | 190.3 | 191.95
|  |  | Passed? (all within 2 mm of 190 mm) | YES | YES
2 | High-contrast spatial resolution | 1 | Best row resolution | 0.9 | 0.9
|  | 1 | Best column resolution | 1 | 0.9
|  |  | Passed? (resolution at least 1 mm) | YES | YES
3 | Slice thickness accuracy | 1 | Ramp length (top) | 39.61 | 54.1992
|  | 1 | Ramp length (bottom) | 44.35 | 62.7441
|  | Derived | Slice thickness | 4.1846 | 5.8159
|  |  | Passed? Within 5 mm +- 1 mm | YES | YES
4 | Slice position accuracy | 1 | Bar length difference (mm) | -1.1035 | -3.0566
|  | 11 | Bar length difference (mm) | -3.9551 | -5.9961
|  |  | Passed? < 5 optimal, 5-7 ok | YES | YES
5 | Image intensity uniformity | 7 | Large ROI low-signal value | 0.4245 | 0.3716
|  | 7 | Large ROI high-signal value | 0.7168 | 0.6444
|  | Derived | Percent integral uniformity (PIU) | 74.3889 | 73.1496
|  |  | Passed? > 82% optimal, <80% fail | NO | NO
6 | Percent-signal ghosting | 7 | Large ROI mean pixel intensity | 0.5158 | 0.4566
|  | 7 | Left ROI MPI | 0.005 | 0.005
|  | 7 | Right ROI MPI | 0.005 | 0.0052
|  | 7 | Top ROI MPI | 0.0134 | 0.0047
|  | 7 | Bottom ROI MPI | 0.0173 | 0.0051
|  | Derived | Ghosting ratio | 0.020066 | 0.00043802
|  |  | Passed? < 0.025; fail if > 0.030 | YES | YES
7 | Low-contrast object detectability | 8 | # Complete spokes | 4 | 9
|  | 9 | # Complete spokes | 10 | 10
|  | 10 | # Complete spokes | 10 | 10
|  | 11 | # Complete spokes | 10 | 10
|  | Derived | Total score | 34 | 39
|  |  | Passed? >37 | NO | YES

## Derived or display values
Name | IRSE Pulseq | IRSE Siemens | TSE Pulseq | TSE Siemens
Test 1 display center | 0.25 | 0.19705 | 0.23 | 0.23
Test 1 display width | 0.5 | 0.3941 | 0.46 | 0.46
Test 3 top mean | 0.0657 |  | 0.0653 | 0.04
Test 3 bottom mean | 0.0658 |  | 0.069 | 0.0418
Test 3 mean | 0.06575 | 0.0383 | 0.06715 | 0.0409
0.032875